## Slide 1: Contained Laparoscopic Morcellation System for Gynecological Surgery
Presented By: Michael Barron

## Slide 2: Uncontained Morcellation Procedure
Separate uterus/fibroid from attachments
Cut into small pieces (morcellation)
Pull pieces out of abdomen

## Slide 3: Need for Contained Laparoscopic Morcellation Systems
Uncontained power morcellation can spread undiagnosed cancer
The alternative is contained power morcellation using an isolation bag
Currently, establishing containment with a bag is frustrating and time consuming
A containment device is needed that:
  Can insert the isolation bag through laparoscopic incisions
  Will reduce the hassle and time used to deploy the bag
  Will endure contact with laparoscopic graspers/endoscope during surgery

## Slide 4: Project Scope
The goal is to design a containment system that decreases overall surgery time by 2 min
Currently, establishing containment can take an extra 10min - 30min of surgery time
  Increasing the efficiency of this process by reducing deployment time would reduce surgical costs as well as infection risk to the patient

## Slide 5: Specifications
Bag must be safe for the patients
Bag deployment must not injure the patient
Waterproof and durable
  Needs to withstand contact with laparoscopic graspers and endoscopes
Number, size and placement of incisions will not change
The bag will fit through a 15mm trocar
Appropriate deployment size to fit in the abdominal cavity.
Priced competitively to existing bags
  Current bags: $2 - $30
  Each minute in the OR: $60 - $100
  Maximum bag price= (current bag price) + ($60)*each minute saved in the OR

## Slide 6: Our Design
Our design will address the following 3 issues:
  Bag deployment method
  The containment bag itself
  Creating a “port” for the side trocar to access

## Slide 7: Bag Deployment Designs
One aspect of our design is how to deploy the containment bag
  Needs to place the bag inside the abdomen in a timely manner
    Bag must be able of passing though the 15mm tocar in the umbilical port
  Must not cause any injury to the patient
  Should be relatively easy for the surgeon to use
  Here are some examples:

## Slide 8: Tube with Plunger
3 component system
  Rigid Tube, Bag, and Plunger
Bag is tightly pre-rolled into a cylindrical shape that is located inside the rigid tube
During operation, the front end of the tube is matched with the 15mm trocar in the umbilical port
Surgeon uses plunger to push bag into abdomen

## Slide 9: Trigger Mechanism
Reusable handle with trigger mechanism
Disposable rod and bag
Trigger is attached to wiring, which runs the length of the tube and into the opening rim of the bag
Activating the trigger will push some of the wire, and thus the bag, out of the tube
  This will expand and open the bag
Deactivating the trigger will pull the wire back into the tube, drawing the bag closed.

## Slide 10: Lubricated Bag
Apply biologically compatible lubricant to either the inside of the trocar or the outside of the isolation bag
Roll the bag into a cylindrical shape
Push through the trocar (method current in use)
The reduction in friction between the outside of the bag and the inside surface of the trocar should allow the bag to pass through more easily

## Slide 11: Hollow Ribbing with Wire
2 parts:
  Bag with hollow ribbing (5-10 ribs)
  Reusable wire device
Wires from the wire device insert into the hollow ribbing. This:
  Compresses bag into tight, cylindrical shape
  Provides rigidity to bag
Use wire device to push bag into abdomen through trocar, then pull the wires out, leaving only the bag in the abdomen

## Slide 12: Hollow Ribbing with Air
Similar to above
Provide rigidity to ribbing with compressed air instead of wire
  Single port in rim of bag connects to each rib
  Pumping compressed air into the port fills the ribs, providing rigidity to the bag
  Open port once bag is in the abdomen, letting the air out and making the bag more flexible

## Slide 13: Containment Bag Designs
Another aspect of our design is what features to include in the containment bag itself
  Should allow the surgeon to better manipulate the bag once within the patient’s abdomen
  Must not cause any injury to the patient

## Slide 15: Colored Tabs
Place brightly colored tabs at the opening and corners of the bag
  Corners will be colored differently than the ones at the opening
Allow quick recognition of the location of the bag by the surgeon
Provides additional location for the bag to be maneuvered by the surgeon using a grasper

## Slide 16: Colored Bag
Apply a transparent color to the base of the bag
Could be striped or solid
Color helps identify where the base of the bag is to the surgeon
Transparency allows the surgeon to still see abdominal bleeding

## Slide 17: Attached Plastic Ring
Traditional containment bags contain a synthetic drawstring along the rim to close the containment bag
  Ends of drawstring are tied together
  Surgeon pulls on drawstring with a grabbing tool
    Slow and clumsy
Instead, attach a plastic ring to the ends of the drawstring
  Allows the surgeon to hook the ring, and then pull the bag closed

## Slide 18: Closed Loop Drawstring
Alternatively, have one continuous drawstring woven into the opening rim of the bag
Include multiple openings in the rim as access points
  Grabbing just one of these openings would allow the surgeon to pull the bag closed

## Slide 19: Spring-loaded Bag:
Include a torsion spring (hair-clip) attached to 2+ arms in the bag
Compress arms together
Insert through the trocar
Bag opens inside abdomen
Close bag later with drawstring

## Slide 20: Side Trocar Access Port Designs
The final aspect of our design is how to facilitate side trocar access to the inside of a deployed containment bag
  Should allow the surgeon to insert the side trocar into the containment bag more easily
  Must keep escaped tissue at a minimum

## Slide 22: Pre-existing Ports
Could be either external or internal
Consists of an off-the-shelf pressure seal built into the lining of the bag and a positioning device
The positioning device enables the surgeon to maneuver the pressure seal next to the side trocar. It could be:
  An elevated ring around the pressure seal, to be hooked by the surgeon
  A string attached to the edge of the pressure seal

## Slide 23: Built in Structural Weakness
Instead of including a specialized port, allows the surgeon to more easily puncture the bag with the side trocar (current method)
Sealed with a balloon-tipped cannula
On the side of the bag, slightly weaken the wall of the bag with:
  Broad areas of slightly thinner wall
  Multiple, cross-shaped indent arranged in a grid pattern

## Slide 24: Final Design Analysis
Ultimately, we decided on:
  A Tube with Plunger as our deployment system
  A Closed Loop Drawstring method for closing the bag
    Also will pursue colored tabs and bag
  Structural weakness as our Port for Side Trocar Access

## Slide 25: Bag Deployment Design Summary

## Slide 26: Pugh Chart of Bag Deployment Designs

## Slide 27: Why we chose it: Tube with Plunger
Highest time-saved estimate of all deployment designs
Simple design
  Reliable
Relatively lightweight
Usable

## Slide 28: Containment Bag Design Summary

## Slide 29: Pugh Chart of Containment Bag Designs

## Slide 30: Why we chose it: Closed Loop Drawstring
More accessible
  Leads to less maneuvering
  Thus more time saved
Minimal space occupied

## Slide 31: Ports for Side Trocar Access Design Summary

## Slide 32: Pugh Chart of Ports for Side Trocar Access

## Slide 33: Why we chose it: Structural Weakness
Easy to align with trocar
Little to none maneuvering necessary
No additional tools needed
Easy to find
Less force is needed to insert trocar into bag
Fast

## Slide 34: Preliminary Design Schedule

## Slide 35: Organization of Team Responsibilities
Dana Sprague – Manages public relations and communication between the team and client
Michael Barron – Leads design generation and defining design specifications
Braden Perkins – Generates CAD sketches to aid with conceptual understanding during brainstorming sessions

## Slide 36: References
"Power Morcellation and Occult Malignancy in Gynecologic Surgery." The American College of Obstetricians and Gynecologists (2014): 1-6. Print.
Einarsson, Jon I., et al. "In-Bag Morcellation." Journal of Minimally Invasive Gynecology 21.5 (2014): 951-3. Print.
“Head Scratcher.” Digital image. Bedol What’s Next. FlugelSoft. 2014. Web. 20 Oct. 2014
Wolfe, Morgan. Personal interview. 29 Aug. 2014, 1 Oct. 2014, 20 ct. 2014.
"Laparoscopy Surgery in Gynaecology." Melaka Fertility. Selva's Fertility, Obsterics & Gynaecology Clinic, n.d. Web. 26 Oct. 2014.
"Hair Clip Spring Supplier Torsion Spring Supplier In China Factory - Buy Hair Clip Spring Supplier Torsion Spring Supplier,Adjustable Torsion Springs,Torsion Spring Constant Product on Alibaba.com." Www.alibaba.com. N.p., n.d. Web. 26 Oct. 2014.
"3M™ Steri-Drape™ Isolation Bag 1003." 3M Infection Prevention. 2014.Web.<http://solutions.3m.com/wps/portal/3M/en_US/IPD-NA/3M-Infection-Prevention/products/catalog/~/3M-Steri-Drape-Isolation-Bag-1003?N=4294957758+5604328&rt=d

## Slide 37: Questions?